# Cleaning Procedures for Laboratory Material used for 
Chemical Preparation/Column Separation in Lead Isotope Analysis

Abbreviations:

- MQ water: Ultrapure water („Milli-Q“ water)

- *** = triple-distilled

## Cleaning reagent vessels/vials

- Rinse three times with MQ water

- Fill with 6% HNO3, make sure that no air is trapped in beakers/vials, let it sit for at least 1 night

- Empty HNO3, fill with MQ water, let it sit for at least 1 night

- Store until needed

## Cleaning Teflon (“Savillex”) beakers

### Pre-cleaning

- Place beakers in large beaker and fill it with a mixture of lab detergent (e.g., DECONEX) and deinoised water, let it sit for at least 1 night

- Rinse three times with deionised water

- Fill the large beaker with Aqua Regia (4/8 deionised water, 3/8 HCl p.a., 1/8 HNO3 p.a.), let it sit for at least 24 h on a hot plate at approx. 70–80 °C

- Rinse three times with deionised water

- Rinse once with MQ water

- Fill the large beaker with 6N HCl, let it sit for at least 24 h on a hot plate at approx. 70–80 °C

- Rinse three times with MQ water

- Take beakers out of large beaker and let them dry

### Bleaching

- Fill 1 ml 6M HNO3*** into the individual beakers, screw tightly, leave for at least one week on a hot plate

- Rinse three times with MQ water

- Fill with MQ water and add a few drops of HBr (to bind the lead), leave it for at least one night or until needed

- Dry beakers at least one day before needed (they should be cooled down to room temperature for weighing)

- Store until needed

## Cleaning columns

- Allow columns filled with old resin to dry

- Empty used resin (tap or rinse with water using a pipette repeatedly)

- Rinse with MQ water

- Place in sealable plastic container with a mixture of lab detergent (e.g., DECONEX) and deinoised water (make sure no air is trapped in the columns) and shake closed container

- Rinse three times with MQ water

- Put columns in large beaker filled with Aqua Regia (4/8 deionised water, 3/8 HCl p.a., 1/8 HNO3 p.a.), let it sit for at least 24 h on a hot plate at approx. 70–80 °C

- Rinse three times with MQ water

- Fill the large beaker with 6N HCl, let it sit for at least 24 h on a hot plate at approx. 70–80 °C

- Rinse three times with MQ water

- Take columns out of beaker and store them 5% HCl*** for at least 1 week or until needed

- Rinse with MQ water before use

## Cleaning new resin (dry product)

- Pour some spoons of resin (e.g., DOWEX AG1x8 or other in use) into a 1 L bottle

- Fill up to approx. 500 ml with MQ water and 10 ml 6M HCl***

- Shake several times

- Decant

- Repeat multiple times

- Fill up again with approx. 500 ml of MQ water and 10 ml 6M HCl***

- Store like this until needed

These procedures were adopted from the former isotope laboratory of Friedrich Begemann, Max-Planck-Institut für Chemie, Dept. Isotopenkosmologie, Mainz, Germany.